- SCOPE

This procedure describes the calibration process for high angle hand-held retroreflectometers using secondary reference plates calibrated in accordance with TP992. These devices simulate an on-road geometry of approximately 12 metres.

- REFERENCES

Operating manual for retroreflectometer.

- DEFINITIONS

Retroreflectivity (mcd /m² /lx) - The ratio of the luminance of a surface to the illuminance at the retroreflector on a plane perpendicular to the direction of the incident light. (Coefficient of retroreflected luminance)

Secondary Reference Plates – boxed set of plates for the calibration of working retroreflectometers, calibrated in turn in accordance with procedure TP992.

Measurement Geometry – Entrance and observation angles defined by the European Committee for Standardisation (CEN), the US Federal Highway Administration (FHWA), and ASTM’s E- 1710 standard.

High Angle Geometry – Entrance angle 86.5°, observation angle 1.5°.

- APPARATUS AND MATERIALS

4.1	Apparatus

High Angle Retroreflectometer – Mirolux 7, Mirolux 12 or equivalent geometry.

Secondary Reference Plates – a boxed set of 4 reference plates, comprising ”zero”, “low”, “medium” and “high” retroreflectivity. (RIAA)

  - Materials

Nil.

- SAMPLING AND SAMPLE PREPARATION

5.1	Sampling

Nil.

  - Sample Preparation

Nil.

- PROCEDURE TO BE FOLLOWED

  - Record the identification number of the boxed standard set.

Notes:

- Calibration shall be carried out with the reference plates mounted in the box as received.

- Reference plates must be cleaned using a soft damp cloth, followed by drying using low pressure compressed air, at least every 2 months. Refer to appendix 1.

  - Warm up and stabilise the retroreflectometer in accordance with the manufacturer’s operating instructions. Ensure that the instrument is fitted with an effective light shroud.

  - Place instrument in secondary standard box over the black panel, and adjust the reading to the specified value ± 1 unit, using the zero adjustment.

  - Where instrument is used for establishing compliance with a specification, place instrument in secondary standard box over the panel nearest to the specified compliance value and adjust to the calibration figure ± 1 unit, using the calibration adjustment. Where adjustment is required, re-check the zero and repeat calibration until both the zero and the calibration are correct.

  - Calibration shall be repeated at intervals of not more than 15 minutes.

  - Where instrument is used not for compliance, but for the determination of a broad range of values, the instrument should be zeroed in accordance with 6.3 above. Calibration shall be carried out using all panel in the set, and a calibration graph developed as per the form attached.

- CALCULATIONS

Nil

- PRECISION

An uncertainty of measurement of ± 7.0% has been determined for this method, at a confidence level of 95% with a nominal coverage factor of 2.

Uncertainty of Measurement determined by and for DPTI Technical Services Group, Field Services Section.

- DOCUMENTATION

Nil

Secondary Reference Set Identification Number:

Instrument Type:

Date Calibrated (office):

Calibration Chart:

Comments:

This test chart is valid for up to 3 months from the date of test.

| Signed |  |
| --- | --- |
| Print Name |  |
| Title |  |
| Date |  |

Form TP994-1, Revn.2.3

Appendix 1

Cleaning of Primary and Secondary Reference Plates

- PURPOSE

The procedure describes the processes to be used when cleaning retroreflective reference panels.

- SCOPE

This procedure applies to reference panels consisting of paint and glass beads applied over a glass substrate.

3.0	     MATERIALS AND EQUIPMENT

- Plastic container, minimum capacity 8 L

- 50mm soft synthetic bristle paint brush

- Dry compressed air supply

4.0	     PROCEDURE

4.1	Place a minimum of 4 L of warm (25-35°C) fresh water in the plastic container.

4.2	Hold the panel by the edges only, wet the brush in the water and gently pass the brush from the top to the bottom of the panel, with the bristles trailing at an angle of approximately 45° to the panel.

4.3	Dip the brush into the water, and repeat the brushing action over the adjoining section of plate, repeating until all of the panel has been covered.

4.4	Rinse the panel surface under flowing fresh water, then allow to drain for a few seconds.

4.5	Dry the panel surface using compressed air.

4.6	Panel shall be tested immediately after drying.

Note: The face of panels must never be touched by bare hands as contamination may occur.